## Slide 1
SU0113A2
SPJSGU8NNNNG
Product brief
Product Description
5V,7W heating pad
Fabric(Suede)
Printed PET film heater
3 step LED controller
Features and Benefits
Uniformity
Thin thickness
Flexibility
Key Applications
Heating pad
Warmer

## Slide 2
Table of Contents
Index | page
Product Brief | 1
Table of Contents | 2
Performance Characteristics | 3
Characteristic Diagram | 4
Reliability Test | 5
Mechanical Dimensions | 6
Product Nomenclature | 7
Packaging | 8
Precaution For Use | 9
Product Data Sheet
SU0113A2/ SPJSGU7NNNNG

## Slide 3
Performance Characteristic
* Product is not guaranteed for operations above the above conditions
Absolute Maximum Rating
Category | standard | Tolerance | Remarks
product name | SU0113A2 |  | 
P/N | SPJSGU7NNNNG |  | 
Fabric | Suede, 0.7t | ±0.3mm | 
Size | Abdomen Band: 1130mm X 185mm X 0.5mm Mat: 250mm X 130mm X 1.5mm Controller: 25mm X 35mm X 10mm Portable Power Bank: 64mm X 93mm X 13mm | ±2mm | 
Weight | Abdomen Band: 140g Mat: 30g Controller: 20g Portable Power Bank: 88g | ±10% | 
Rated voltage | 5 DCV | - | 
Resistance | 3.5Ω | ±10% | 
Electricity | 7W | ±10% | 
Power density | 0.19W/cm2 | ±10% | 
Temperature of use | LEVEL : 1 (40℃) 2 (45℃) 3 (50℃) |  | Sensor control
Operating condition 1 | MAX. 90min |  | off
Operating condition 2 | Insulation thermal 70℃ |  | off
* Resistance is a measurement tolerance
* Temperature of use is single product operating environment
* After turn on the power button, power on and set to level 3 automatically
 | Voltage used | Heater temperature | Operating conditions
Single Product Operating Environment | ~ 7V | 80℃ | at 25℃
Product Data Sheet
SU0113A2/ SPJSGU7NNNNG

## Slide 4
Product Data Sheet
ON
50℃
ON
45℃
ON
40℃
LEVEL1
LEVEL2
LEVEL3
Characteristic Diagram
Test method : 25℃, 1 Measurement of saturation temperature at voltage rise
Chart: TEMPERTATURE(℃)
| LEVEL1 | LEVEL2 | LEVEL3
 | 41.4 | 47.1 | 52.2
SU0113A2/ SPJSGU7NNNNG

## Slide 5
Item | Test Condition | Unit | Number of Damage | Judgment | Result
Folded test | Folded, 1 cm in diameter, 80 kg Press repeatedly and measure current, resistance check | 4,500 Cycle | 0/3 | <±8% | OK
Twist test | 180° twist , resistance check (End to end :1kg force) | 15,000 Cycle | 0/3 | <±8% | OK
Press hard test | 10 cm in diameter, 80 kgf, resistance check | 60,000 cycle | 0/3 | <±8% | OK
continuous operating | Rated 5V, resistance check | 2,000 Hours | 0/3 | <±8% | OK
Reliability Test
※ Judgement is change before and after
Product Data Sheet
SU0113A2/ SPJSGU7NNNNG

## Slide 6
Mechanical Dimensions
Product Data Sheet
Notes : All dimensions are in mm
SU0113A2/ SPJSGU7NNNNG

## Slide 7
P1 | P2 | P3 | P4 | P5 | P6 | P7 | P8 | P9 | P10 | P11
Series | product | Material | Color | Power type | W | the closing part | Thermoelectronic Film | Remote | Labels | Cover
S-Soft On | P - USB Play | S- Suede | R-Rosewood | U-USB | 5 | N-Normal (Gray) | N-Normal (Black) | N-Normal (Gray) | N-Normal | N- X
 | H - portable heating cushion | T- TPU | B-Blue | D1-DC 12V | 7 | CC-Customized Color | C-Customized color & logo | NL-Normal LOGO (INKO) | CC-Customized Color | G-Gray
 | J - Thermal heater | M- Micro Hiber | G-gray | D2-DC 24V | 9 | CL-Customized Logo |  | CC-Customized Color | CL-Customized Logo | R-Rosewood
 | JN - Slim Fit Thermal heater | P-Poly ester | BE-beige |  | 37 |  |  | CLO-Customized Logo | C-Color & Logo both | CC-Customized Color
 | PJ- Pocket-type thermal Heater |  | N-Navy |  | 60 |  |  | CLI-Customized Line (Change color/length) |  | CL- Customized Logo
 | S- sleep+ a portable heating mat |  | FG-Forest Green |  |  |  |  |  |  | 
 | B- NEW premium microfiber heating blanket |  | Y-Yellow |  |  |  |  |  |  | 
 | U- slim Detachable USB Hot pack |  | BR-Amber Brown |  |  |  |  |  |  | 
 | C- Memory foam heating cushion |  | SP-Soft pink |  |  |  |  |  |  | 
Product Nomenclature
① Part Numbering System
② Serial Numbering System
Product Data Sheet
P1P2P3P4P5P6P7P8P9P10 SPJSGU7NNNNG
For quality and LOT tracking
Management, attach QR barcode
Stickers to the back of the controller.
Model Information |  |  |  |  | Manufacturing Information(LOT Number) |  |  |  |  |  |  | Assembly Date |  |  |  |  | Serial number
Material | Marking | Company /APP | Marking |  | Year of production | Marking | Month | Marking | Date | LOT order |  | Year of production | Marking | Month | Marking |  | Sequence
PET | FE | client 1 | 00 |  | 2023 | 3 | 1 | A | 00 | 00 |  | 2023 | 3 | 1 | A |  | 0000
PI | FI | client 2 | 01 |  | 2024 | 4 | 2 | B | 01 | 01 |  | 2024 | 4 | 2 | B |  | 0001
SUS | US | client 3 | 05 |  | 2025 | 5 | 3 | C | 02 | 02 |  | 2025 | 5 | 3 | C |  | 0002
Al | UA | client 4 | 02 |  | 2026 | 6 | 4 | D | 03 | 03 |  | 2026 | 6 | 4 | D |  | 0003
USB | SU | client 5 | 03 |  | 2027 | 7 | 5 | E | 04 | 04 |  | 2027 | 7 | 5 | E |  | 0004
adaptor | SD | client 6 | 04 |  | 2028 | 8 | 6 | F | 05 | 05 |  | 2028 | 8 | 6 | F |  | 0005
 |  | . | . |  |  |  | 7 | G | 06 | 06 |  |  |  | 7 | G |  | 0006
 |  | . | . |  |  |  | 8 | H | 07 | 07 |  |  |  | 8 | H |  | 0007
 |  | . | . |  |  |  | 9 | I | 08 | 08 |  |  |  | 9 | I |  | 0008
 |  |  |  |  |  |  | 10 | J | 09 | 09 |  |  |  | 10 | J |  | 0009
 |  |  |  |  |  |  | 11 | K | ~ | ~ |  |  |  | 11 | K |  | ~
 |  |  |  |  |  |  | 12 | L | 31 | 99 |  |  |  | 12 | L |  | 9999
SU0113A2/ SPJSGU7NNNNG

## Slide 8
Packaging
Packing Box
Product Data Sheet
SU0113A2/ SPJSGU7NNNNG
Components
Packing Box

## Slide 9
Precaution for Use
Use carefully to avoid the risk of low-temperature burns if used for a long time.
Do not touch with a heat-generating product.
(Do not use for infants and toddlers under 3 years old.)
Do not use this device on the elderly who are sensitive to heat or cannot cope with
overheating.
If you have sensitive skin, it is recommended to use after a test or consultation.
If the product is damaged or broken, do not plug it in and contact the customer
service center.
Do not damage or arbitrarily disassemble, repair, or modify with sharp instruments
Keep the product away from fire.
Avoid wet areas when using or storing, and do not connect to power, especially with
wet hands.
Do not fold or crumple severely when using or storing.
Product Data Sheet
SU0113A2/ SPJSGU7NNNNG

## Slide 10
HEAD OFFICE
10, Sandan 4-gil, Seo-myeon, Suncheon-si, Jeollanam-do, Republic of Korea,(57927)
FACTORY
49, Hangman 10-ro, Gwangyang-si, Jeollanam-do, Republic of Korea(57793)
SEOUL OFFICE
PARU B/D 71, Mokdongjungangbuk-ro, Yangcheon-gu, Seoul Korea(07947)
TEL +82-61-755-5114
FAX +82-61-753-0134